## Slide 1: Instrument Tests with the new Leica AT401
Georg Gassner September 16th 2010

## Slide 2: Leica Specifications
Distance:
  Range: ~1.5 m - ~80 m (Brochure)
  Accuracy: < 10 µm (Factsheet)
  MPE: ± 10 µm (Brochure)
Angle:
  Vertical Range 0° - 145°
  0.5 arcsec (2.5 µrad)
  MPE: ± 15 µm + 6 µm/m
UXYZ = ± 15 µm + 6 µm/m (±45°)
Level to gravity
PowerLock
Battery Operated
CCD Camera
Notes: Added – to range, spaces between values and units, ° symbol to angles in degrees

## Slide 3: Distance Measurements - Setup
Test Bench at SLAC in Sector 10
Notes: Changed “Sect 10” to “Sector 10”

## Slide 4: Distance Measurements - Different Measurement Time Settings

## Slide 5: Distance Measurements - Different Measurement Time Settings - Zoom
Std: 1.5 µm
Laboratory environment
Distances 1.5 m to 32 m

## Slide 6: Distance Measurements – Small Sampling Distance

## Slide 7: Horizontal Angle Test - Setup
10 setups with the instrument at the same position but rotated by 30° Targets at same height as instrument
  Horizontal Angle Meas.
  Manual Targeting
  Level of Instrument
Notes: Shockingly, removed period after 30° since there was none after “instrument”

## Slide 8: Horizontal Angle Test – Hz-Results
Std: 1.6 µrad (0.3 arcsec)
Laboratory environment
Distances 3 m to 8 m

## Slide 9: Horizontal Angle Test – Level Sensor
Roll [µrad]
Pitch [µrad]

## Slide 10: Warm-up Test – Setup
Set measurement every 5 minutes after the instrument was turned on

## Slide 11: Warm-up Test – Results (1)

## Slide 12: Warm-up Test – Results (2)
Repeatability [µrad] resp. [µm]

## Slide 13: Acceptance/Verification Test – ISO 10360-LT
6.4 Length Error
  Measure a multitude of 3D distances in space and compare them to a “known” distance. Compare the difference to the manufacturer’s MPE specification or your requirements.

## Slide 14: Acceptance/Verification Test – Example
3m
10 m
Notes: “Example of a realization”, sounds redundant, whatever a “realization” is. “An Example” would be sufficient. It’s up to you.

## Slide 15: Acceptance/Verification Test – Results

## Slide 16: Summary
Pro
  Portable
  Accurate
  PowerLock
  Level sensor
  CCD camera
  Long range
Con
  Official min. dist 1.5m
  No scanning
  PowerLock (new firmware?)

## Slide 17: Close Range (<1.5 m)