## Slide 1: Hydrostatic Level Systems at Fermilab and DUSEL
J Volk, V Shiltsev
Fermilab USA
A Chuprya, M Kondaurov, S Singatulin
Budker Institute of Nuclear Physicis Russia
L Stetler, J Van Beek
South Dakota School of Mines and Technology USA

## Slide 2: Hydro static water Level Systems HLS
BUDKER sensor
Capacitive pickup
Accuracy 1 micrometer
Cost $1200 per channel
Capacitive sensor
Water pool
On stand with water and
Air line connections
Water line
Air line

## Slide 3: Ultra Sonic Sensor and Electronics
Water pool and sensor
Separate electronics
Ultra sonic sensor better than 1 micrometer resolution $4000 per channel

## Slide 4: Schematic of Ultra Sonic Sensor
R1 and R2 are
Fixed distances
used for calibration
OF is water level
Target at top is
for alignment

## Slide 5: Ultra Sound Pulses
Description | Min | Actual | Max | units
Echo pulse duration | 0.2 | 0.8 | 2 | μs
Test Frequency | 5 | 6.9 | 10 | MHz
Bandwidth at 6 db | 15 | 25.9 | 50 | %
Pulse echo sensitivity | -48 | -42.2 | -38 | dB
2 volts per division
20 μs per division
GE H10KB5T

## Slide 6: Fermilab design Tevatron style
8 channel readout card
Ether net interface
Balluff sensor and pool
Power supply
RS232 port
Paper in IWAA-08 J Volk

## Slide 7: Sensors at Fermilab and DUSEL
Name | Location | Type of Sensor | Number of sensors | Software version | Operating system
Tev Sensors | Tevatron quadrupoles | Balluff | 204 |  | ACNET
B0 | Low beta quads B0 interaction region | SAS | 9 | 19_Mar-09 | Windows XP
D0 | Low beta quads D0 interaction region | SAS | 9 | 19_Mar-09 | Windows XP
HLS-Minos-2 | MINOS Hall | SAS-E | 7 | 04-Mar-10 | Windows XP
ADAPS90581 | LaFarge Mine North Aurora Il | SAS | 5 | 19_Mar-09 | Windows XP
Array A | 2000-ft level Homestake Gold mine, Lead SD | TEV | 6 |  | Windows XP
Array B | 2000-ft level Homestake Gold mine, Lead SD | TEV | 6 |  | Windows XP
Array C | 2000 ft level Homestake Gold mine Lead SD | SAS-E | 6 | 04-Mar-10 | Windows XP
DUSEL-1 | MP8 tunnel Fermilab | SAS-E | 12 | 04-Mar-10 | Windows XP
HLS-MP8-1 | MP7 tunnel Fermilab | ULSE | 3 | 04-Mar-10 | Windows XP
HLS-MP8-2 | MP8 tunnel Fermilab | ULSE | 12 | 04-Mar-10 | Windows XP
BDPHY14 | MP8 tunnel Fermilab |  | Test and calibration | 04-Mar-10 | Windows 7
Hot Spares | MP8 tunnel Fermilab | SAS-E | 4 | 04-Mar-10 | Windows 7
PK 50 | Porta Kamp 50 Fermilab |  | Test and calibration | 04-Mar-10 | Windows XP

## Slide 8: Tevatron HLS sensors during quench
Chart: Date

## Slide 9: Quench Low Beta Quads B0 Interaction Region
A side where quench occurred
B side response to quench

## Slide 10: One Month MINOS Data

## Slide 11: Other types of Ground Motion
Subsidence caused by earth quakes
Floor tilting caused
by movement of
water
Primary
Secondary

## Slide 12: Five Years of MINOS data

## Slide 13: Two Years of Data From LaFarge Mine

## Slide 14: Chilean Earth Quake of February 27, 2010
8.8 magnitude earthquake as seen
By the LaFarge mine HLS
Secondary wave tilted floor
For 25 minutes
16 μm
Spacing 120 m
0.13 μradian

## Slide 15: DUSEL Lead South Dakota
Homestake gold mine
Cross section of Homestake gold mine
Ross shaft
Yates Shaft
1.2 km

## Slide 16: HLS in 2000 ft Level Homestake Gold Mine DUSEL/Sandford lab

## Slide 17: Array C Budker Sensors 2000 ft level Homestake Gold mine
Difference in two sensors 2000 ft level

## Slide 18: Calibration Stand for Balluff sensors
Calibration of a Balluff sensor
Adjustable stage
and micrometer

## Slide 19: Balluff Sensor Calibration Data

## Slide 20: View of SAS-E Test Stand
Single pipe half filled system
Pipe diameter 25.4 mm

## Slide 21: Two sensors on same system 4 m apart 2 weeks of data
Chart: Date

## Slide 22: Adding 50 CC of water to system
Chart: Date

## Slide 23: Adding 50 cc of water to SAS-E system
 | SN 188 | SN 189 | SN 190 | SN 191 | SN 193 | SN 194 | SN 195 | SN 196 | SN 197 | SN 198 | SN 199 | SN 200 | Average | Std Dev
Fill number | microns | microns | microns | microns | microns | microns | microns | microns | microns | microns | microns | microns | microns | microns
1 | 232.7 | 232.4 | 232.4 | 232.3 | 232.3 | 232.3 | 232.2 | 232.0 | 232.1 | 232.1 | 231.7 | 232.0 | 232.2 | 0.3
2 | 236.5 | 236.6 | 236.2 | 236.4 | 236.6 | 236.8 | 236.6 | 236.4 | 236.5 | 236.2 | 235.4 | 236.0 | 236.3 | 0.4
3 | 238.7 | 238.7 | 238.6 | 238.5 | 238.6 | 238.8 | 238.7 | 238.6 | 238.6 | 238.5 | 237.8 | 238.1 | 238.5 | 0.3
4 | 239.5 | 239.4 | 239.3 | 239.1 | 239.2 | 239.4 | 239.1 | 239.2 | 239.2 | 239.1 | 238.2 | 238.7 | 239.1 | 0.3
5 | 241.3 | 241.2 | 241.0 | 240.7 | 241.0 | 241.1 | 241.0 | 241.0 | 241.0 | 240.5 | 240.3 | 240.5 | 240.9 | 0.3
6 | 233.8 | 233.6 | 233.6 | 233.6 | 233.6 | 233.6 | 233.6 | 233.7 | 233.7 | 233.6 | 232.9 | 233.3 | 233.6 | 0.2
7 | 243.9 | 243.9 | 243.7 | 243.6 | 243.6 | 243.8 | 243.8 | 244.0 | 243.8 | 243.6 | 243.0 | 243.3 | 243.7 | 0.3
8 | 246.6 | 246.3 | 246.4 | 246.2 | 246.3 | 246.1 | 246.3 | 246.4 | 246.4 | 246.3 | 245.5 | 245.9 | 246.2 | 0.3
9 | 237.9 | 237.4 | 237.6 | 237.4 | 237.6 | 237.5 | 237.7 | 237.7 | 237.7 | 237.4 | 236.8 | 237.0 | 237.5 | 0.3
10 | 247.0 | 246.7 | 246.9 | 246.8 | 246.8 | 246.8 | 246.9 | 247.0 | 246.9 | 246.8 | 246.4 | 246.4 | 246.8 | 0.2
11 | 243.7 | 243.3 | 243.3 | 243.2 | 243.3 | 243.2 | 243.4 | 243.3 | 243.1 | 242.8 | 242.5 | 242.5 | 243.1 | 0.4
12 | 245.2 | 244.9 | 245.1 | 245.1 | 245.0 | 244.9 | 245.1 | 245.2 | 245.1 | 245.1 | 244.8 | 245.0 | 245.0 | 0.1
13 | 222.1 | 222.0 | 222.1 | 222.2 | 221.9 | 221.9 | 222.0 | 222.0 | 221.9 | 222.0 | 221.5 | 221.7 | 222.0 | 0.2
14 | 223.7 | 223.5 | 223.6 | 223.6 | 223.4 | 223.4 | 223.4 | 223.6 | 223.3 | 223.3 | 222.6 | 222.9 | 223.4 | 0.3
15 | 213.6 | 213.6 | 213.7 | 213.9 | 213.6 | 213.7 | 213.9 | 213.9 | 213.7 | 213.8 | 213.3 | 213.5 | 213.7 | 0.2
16 | 226.5 | 226.5 | 226.6 | 226.8 | 226.6 | 226.6 | 226.7 | 226.8 | 226.6 | 226.6 | 226.2 | 226.2 | 226.6 | 0.2
17 | 215.4 | 215.0 | 215.1 | 215.4 | 215.1 | 215.3 | 215.8 | 216.1 | 216.1 | 216.4 | 216.2 | 216.5 | 215.7 | 0.5
18 | 226.8 | 226.7 | 226.7 | 227.0 | 226.6 | 226.7 | 226.9 | 227.0 | 226.8 | 227.2 | 227.0 | 227.1 | 226.9 | 0.2
19 | 229.3 | 229.0 | 229.0 | 229.4 | 228.8 | 229.0 | 229.0 | 229.1 | 228.6 | 228.8 | 228.3 | 228.1 | 228.9 | 0.4

## Slide 24: Adding 50 cc of Water to ULSE Sensors
Fill number | SN 53 | SN 57 | SN 64 | SN 56 | SN 63 | SN 62 | SN 54 | SN 55 | SN 61 | SN 59 | SN 58 | SN 60 | average | std dev
1 | 231.5 | 228.9 | 227.2 | 228.1 | 229.7 | 227.8 | 225.2 | 226.7 | 225.3 | 220.2 | 222.4 | 211.4 | 225.4 | 5.4
2 | 213.8 | 218.7 | 214.7 | 216.7 | 217.2 | 216.7 | 217.5 | 218.8 | 216.5 | 216.5 | 214.2 | 214.1 | 216.3 | 1.7
3 | 221.2 | 221.4 | 222.7 | 223.3 | 221.1 | 220.7 | 222.4 | 223.6 | 221.6 | 220.6 | 218.6 | 218.7 | 221.3 | 1.6
4 | 239.5 | 216.4 | 213.4 | 213.4 | 211.2 | 209.5 | 210.6 | 212.8 | 211.6 | 219.3 | 209.0 | 227.7 | 216.2 | 9.0
5 | 226.8 | 219.6 | 217.1 | 216.2 | 213.8 | 211.9 | 214.0 | 214.9 | 213.4 | 217.4 | 211.5 | 220.9 | 216.5 | 4.4
6 | 228.3 | 228.3 | 226.4 | 228.6 | 227.8 | 226.4 | 228.3 | 229.3 | 228.0 | 226.8 | 225.1 | 227.7 | 227.6 | 1.2
7 | 229.4 | 231.5 | 229.9 | 232.0 | 230.7 | 229.2 | 231.3 | 232.1 | 230.7 | 230.6 | 227.5 | 232.2 | 230.6 | 1.4
8 | 232.0 | 229.9 | 228.7 | 230.1 | 228.9 | 227.0 | 229.0 | 229.7 | 228.1 | 228.8 | 225.1 | 229.7 | 228.9 | 1.7
9 | 239.7 | 240.9 | 238.9 | 240.9 | 239.1 | 237.5 | 239.5 | 240.5 | 239.3 | 240.0 | 236.7 | 240.1 | 239.4 | 1.3
10 | 228.0 | 227.7 | 224.9 | 227.2 | 226.0 | 224.9 | 226.7 | 227.3 | 226.1 | 226.5 | 223.5 | 226.1 | 226.2 | 1.3
11 | 244.4 | 244.4 | 242.4 | 244.5 | 242.8 | 241.6 | 243.6 | 244.6 | 243.7 | 244.4 | 241.4 | 244.4 | 243.5 | 1.2
12 | 247.4 | 246.6 | 244.0 | 245.6 | 243.4 | 242.1 | 244.0 | 245.1 | 244.6 | 244.8 | 242.9 | 245.8 | 244.7 | 1.5
13 | 244.8 | 243.3 | 241.0 | 243.2 | 241.6 | 240.8 | 242.5 | 243.7 | 242.8 | 242.1 | 240.8 | 243.3 | 242.5 | 1.3
14 | 233.5 | 236.8 | 234.9 | 236.9 | 235.1 | 234.4 | 235.9 | 237.0 | 236.1 | 241.9 | 247.1 | 237.2 | 237.2 | 3.7

## Slide 25: Heating of Budker Capacitive Sensors with fixed disk in pool
0.5 micro meter per division
SN 198
SN 188

## Slide 26: Heating of Budker Ultra Sonic sensors
Same scale 1 micro meter per division
Sensor 53
Sensor 60

## Slide 27: Heating of electronics only for Budker Ultrasonic sensor
0.5 micro meter per division

## Slide 28: Findings From Calibration Studies
The capacitive sensors calibration is good to 0.5 micro meters as determined by water filling tests.
The temperature dependence of the capacitive sensors has a hysteresis and varies from sensor to sensor.
This could be an electronic effect and will need more testing.
Need to split analog and digital parts test separately.
The ultrasonic sensors have a larger temperature dependence than the capacitive sensors.
The electronics for the ultrasonic sensors is very stable as measured by separate tests.
Change in water level can not be totally accounted for by water expansion.
Need to test transponders to fully understand temperature effects.

## Slide 29: Conclusions
HLS can provide useful information for accelerator operations.
Resolutions of less than 1 micro meter are possible.
Stability test of systems needs more work.
Temperature dependence of electronics needs to be studied.
There is a need to cross calibrate different systems

## Slide 30
Danke schön!
Thank You
For you attention
I can be reached at volk@fnal.gov